|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Product Name** | [Minuet](#Minuet)[http://www.arjohuntleigh.com.au/PageFiles/1448/minuet.jpg](http://www.google.com.au/url?sa=i&rct=j&q=&esrc=s&frm=1&source=images&cd=&cad=rja&uact=8&ved=0CAcQjRxqFQoTCK-jlN-4pccCFQSepgodpYwGsg&url=http://www.arjohuntleigh.com.au/products/medical-beds/community-beds/minuet-community-bed-range/&ei=xzrMVe-jGIS8mgWlmZqQCw&bvm=bv.99804247,d.dGY&psig=AFQjCNHLUOvn0gAJzwfLueOvyyABf9VUvQ&ust=1439534089583785) | [Arminia II](#Arminia) [https://encrypted-tbn3.gstatic.com/images?q=tbn:ANd9GcTu9CI411GtPJnghBF386XBvFlQky-60ZG8V-KnLZ041Ih1yXO4Rw](https://www.google.com.au/url?sa=i&rct=j&q=&esrc=s&frm=1&source=images&cd=&ved=0CAcQjRxqFQoTCPeg_vuypccCFQRQpgodtOsEHg&url=https://www.burmeier.com/en/homecare/special-beds/dali-low-entry&ei=uTTMVbfgDoSgmQW015PwAQ&bvm=bv.99804247,d.dGY&psig=AFQjCNEUvMgW6x02QneRarxl4qkCmBLj1Q&ust=1439532591095530) | MelanieVerstellung-Melanie | [Etude Plus](#Etude)  [http://www.gmsrehab.com.au/images/ETPLUSTHUMB.jpg](http://www.gmsrehab.com.au/images/ETPLUS.jpg) |
| Manufacturer | Arjo Huntleigh | Burmeir | Betten Malsch | Invacare |
| Availability through DES | **refurbished only** | **refurbished only** | **purchase\* or refurbished** | **purchase\* or refurbished** |
| DES Code | AB01I A | AB01O A | AB01P A | AB01T A |
| SWL capacity  NB: allow approx. 40kg for mattress & bedding | 180kg | 175kg | 175kg | 200kg |
| Features | head raise,  knee bend, manual leg raise Trendelenburg reverse Trendelenburg | head raise,  knee bend, leg raise manual,  Trendelenburg  (some have reverse Trendelenburg) | head raise,  knee bend, leg raise manual,  Trendelenburg reverse Trendelenburg | head raise,  knee bend, leg raise manual,  Trendelenburg reverse Trendelenburg |
| Total bed footprint (W x L) | 920 x 2200 | 1030 x 2200 | 1050 x 2130 | 1020 x 2200 |
| Height range | 220 - 630 | 400 - 800 | 400 - 800 | (standard) 410 – 810  OR 340 -740 |
| Clearance under bed | clear | clear | clear | clear |
| Backrest platform length | 760 | 750 | 700 | 850 |
| Backrest angle | 0° - 73° | 0° - 70° |  | 0° - 70° |
| Internal platform dimensions (W x L) | 920 x 2040 | 890 x 2050 | 890 x 2000 | 890 x 2080 |
| Lockable handle controller? | no | no | yes | optional |
| Mattress retainers? | Side (head end) Foot end | Sides | no | yes, can rotate away |
| Width between retainers | 870 (width) 2000 (length) | 850 | n/a | 850 |
| Bed extension available**?** | no | yes – 150mm | no | yes – 200mm |
| Self-help pole position | 90mm in from headboard 70mm from side railing | 75mm in from headboard 70mm from side railing | n/a | 80mm in from headboard  60mm from side railing |
| Other notes |  |  |  | Not suitable for <12yo |

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **Minuet** | | | | | | | | |
| **Bed Rail Options and Specifications** | | | | | | | | |
| Style | 3/4 length | | | | | **C:\temp\ViewingImage.jpg** | | |
| DES Code | AB09G A | | | | |
| Internal gaps within rails | 120 | | | | |
| Distance between left and right bed rails | 980 | | | | |
| Rail height from bed platform | 395 | | | | |
| Rail cover options | Padded – ABO9I A | | | | | C:\temp\ViewingImage.jpg | | |
| **Information on Mattress Requirements** | | | |  | | | | |
| **Configuration** | **Max. Height** | | **Max. Width**  (no side retainers) | | | **Max. Width**  (side retainers) | | **Max. Length** |
| no rails | n/a | | 920 | | | 870 | | 2000 |
| 3/4 rails | 175 | | 940 | | | 870 | | 2000 |
| **Potential Suitable Mattresses** | | **NB: does not indicate entrapment zone compliance** | | | | | | |
| **Product Name** | | **DES Code** | | | **Equip. Program Category** | | **Mattress Dimensions**  **(W x L x H)** | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | | AB14J A | | | Cat 1 | | 870 x 1960 X 150 | ¾ rails |
| Zenith PM265 Mattress Overlay | | AB15AA A | | | Cat 1 | | 900 x 2000 x 127 (+100 mm underlay) | no rails |
| Zenith PM268 Mattress Replacement System | | AB15AB A | | | Cat 1 | | 900 x 2000 x 200 | no rails |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Arminia II** | | | | | | |
| **Bed Rail Options and Specifications** | | | | | | |
| Style | Full length | | | |  | |
| DES Code | n/a (ordered with bed) | | | |
| Internal gaps within rails | 150 | | | |
| Distance between left and right bed rails | 940 | | | |
| Rail height from bed platform | 385 | | | |
| Rail cover options | full length - AB09N A | | | | C:\temp\ViewingImage.jpg | |
| **Information on Mattress Requirements** | | |  | | | |
| **Configuration** | **Max. Height** | **Max. Width**  (no side retainers) | | | **Max. Width**  (side retainers) | **Max. Length** |
| no rails | n/a | 900 | | | 850 | 2010 |
| Full length rails | 165 | 900 | | | 850 | 2010 |
| **Potential Suitable Mattresses** | | **NB: does not indicate entrapment zone compliance** | | | | |
| **Product Name** | | **DES Code** | | **Equip. Program Category** | **Mattress Dimensions**  (W x L x H) | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | | AB14J A | | Cat 1 | 870 x 1960 X 150 | Full length |
| Mattress ROHO 4 Sections Includes Cover | | AB15F A | | Cat 1 | 860 x 2030 x 80 | Full length |
| Zenith PM265 Mattress Overlay | | AB15AA A | | Cat 1 | 900 x 2000 x 127 (+75mm underlay) | Full length |
| Zenith PM268 Mattress Replacement System | | AB15AB A | | Cat 1 | 900 x 2000 x 200 | Full length |

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| **Etude Plus** | | | | | | | |
| **Bed Rail Options and Specifications** | | | | | | | |
|  | **Scala Basic 2** | | | **Full length** | | | |
| **C:\temp\ViewingImage.jpg** | | | **C:\Users\larthu01\Desktop\ESIS images\20150820_143524.jpg** | | | |
| DES Code | AB09J A | | | AB09F A  AB09Q A (extended length rails) | | | |
| Internal gaps within rails | 100 | | | 100 | | | |
| Distance between left and right bed rails | 910 | | | 955 - 970 | | | |
| Rail height from bed platform | 390 | | | platform at lowest setting: 440  platform at highest setting: 370 | | | |
| Rail cover options | C:\Users\larthu01\Desktop\ESIS images\20150820_150551.jpg Padded – AB09K A | | | C:\Users\larthu01\Desktop\ESIS images\20150820_143713.jpgMesh - AB09O A; AB09S A (for extended length rails) | | | |
|  | | | Padded - AB09P A; AB09R A (for extended length rails)  C:\Users\larthu01\Desktop\ESIS images\20150820_141651.jpg | | | |
| **Information on Mattress Requirements** | | |  | | | | |
| Configuration | **Max. Height** | | **Max. Width**  (no side retainers) | | | **Max. Width**  (with side retainers) | **Max. Length** |
| no rails | n/a | | 900 | | | 850 | 2040 |
| Scala Basic 2 | 170 | | 880 | | | 850 | 2040 |
| Full length rails | lowest setting: 220  highest setting: 150 | | 900 | | | 850 | 2040 |
| **Potential Suitable Mattresses** | | **NB: does not indicate entrapment zone compliance** | | | | | |
| **Product Name** | | **DES Code** | **Equip. Program Category** | | **Mattress Dimensions**  (W x L x H) | | **Rail Options** |
| Mattress Hospital Bed MSS Premier Glide | | AB14J A | Cat 1 | | 870 x 1960 X 150 | | all |
| Zenith foam mattress - narrow | | AB14O A | Cat 1 | | 840 x 2020 x 150 | | all |
| Mattress Platinum Foam | | AB14Q A | Cat 1 | | 880 x 2020/2040 x 150 | | all |
| Mattress ROHO 4 Sections Includes Cover | | AB15F A | Cat 1 | | 860 x 2030 x 80 | | all |
| Zenith PM265 Mattress Overlay | | AB15AA A | Cat 1 | | 900 x 2000 x 127 (+75mm underlay) | | Full length |
| Zenith PM268 Mattress Replacement System | | AB15AB A | Cat 1 | | 900 x 2000 x 200 | | Full length |
| Mattress Overlay AlphaXcell | | AB15T A | Cat 2 | | 840 x 2050 x 120 | | no rails |